## MainExperiment_Blind
Exp. | House | [6 empty] | Sensor |  |  |  |  | Data aquisition system |  |  | Channel lists
blind | wet | Test_ (O5) | Ref_ (N2) | cannel name | unit | location | comment | manufacturer | type | calibrator | calibartion date | specified accuracy | specified accuracy | type | manufacturer | Test House (O5) | Calibration/SerialNo. | Reference House (N2) | Calibration certificate
x | x |  |  | AmbientAirTemperature | °C | weather station | operated by Meteomedia GmBH | Lambrecht |  |  | 2010 | 0.1
x | x |  |  | CO2 | ppm | weather station | operated by Meteomedia GmBH |  |  |  | 2010
x | x |  |  | Pressure_Atmosphere | hPa | weather station | operated by Meteomedia GmBH
x | x |  |  | Radiation_Diffuse | W/m² | weather station | operated by Meteomedia GmBH | Kipp & Zonen | CMP11 |  | 2010 | < 2 %
x | x |  |  | Radiation_East | W/m² | weather station | operated by Meteomedia GmBH | Kipp & Zonen | CMP11 |  | 2018 | < 2 %
x | x |  |  | Radiation_Global | W/m² | weather station | operated by Meteomedia GmBH | Kipp & Zonen | CMP11 |  | 2010 | < 2 %
x | x |  |  | Radiation_North | W/m² | weather station | operated by Meteomedia GmBH | Kipp & Zonen | CMP11 |  | 2010 | < 2 %
x | x |  |  | Radiation_South | W/m² | weather station | operated by Meteomedia GmBH | Kipp & Zonen | CMP11 |  | 2010 | < 2 %
x | x |  |  | Radiation_West | W/m² | weather station | operated by Meteomedia GmBH | Kipp & Zonen | CMP11 |  | 2010 | < 2 %
x | x |  |  | RadiationIR_global | W/m² | weather station | operated by Meteomedia GmBH | Kipp & Zonen | CGR4 |  | 2010 | ±6W/m² & ±1%
x | x |  |  | RadiationIR_global_InstrumentTemperature | °C | weather station | operated by Meteomedia GmBH
x | x |  |  | RadiationIR_west | W/m² | weather station | operated by Meteomedia GmBH | Kipp & Zonen | CGR4 |  | 2010 | ±6W/m² & ±1%
x | x |  |  | RadiationIR_west_InstrumentTemperature | °C | weather station | operated by Meteomedia GmBH
x | x |  |  | Rain_driving_Free | mm/h | weather station | operated by Meteomedia GmBH
x | x |  |  | Rain_driving_WestFacade | mm/h | weather station | operated by Meteomedia GmBH
x | x |  |  | Rain_Normal | mm/h | weather station | operated by Meteomedia GmBH
x | x |  |  | RelativeHumidity | % | weather station | operated by Meteomedia GmBH | Lambrecht |  |  | 2010 | ±2% rel.H.
x | x |  |  | SnowHight | mm | weather station | operated by Meteomedia GmBH
x | x |  |  | SoilTemperature_000cm | °C | weather station | operated by Meteomedia GmBH | Heraeus Holding GmbH |  |  | 2010 | 0.1
x | x |  |  | SoilTemperature_050cm | °C | weather station | operated by Meteomedia GmBH | Heraeus Holding GmbH |  |  | 2010 | 0.1
x | x |  |  | SoilTemperature_100cm | °C | weather station | operated by Meteomedia GmBH | Heraeus Holding GmbH |  |  | 2010 | 0.1
x | x |  |  | SoilTemperature_200cm | °C | weather station | operated by Meteomedia GmBH | Heraeus Holding GmbH |  |  | 2010 | 0.1
x | x |  |  | WindDirection | ° | weather station | operated by Meteomedia GmBH | Adolf Thies GmbH & Co.KG |  |  | 2018 | 0...5 m/s: ±0,1m/s of measuring value >5 m/s: ±2% of measuring value
x | x |  |  | WindSpeed | m/s | weather station | operated by Meteomedia GmBH | Adolf Thies GmbH & Co.KG |  |  | 2018 | ±1°
|  | x | x | aroom_bath_010_AT | °C | Bath air temperature @10 cm hight | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_bath_010_AT | WS201809131_19 | N2_aroom_bath_010_AT | WS201809127_16
|  | x | x | aroom_bath_-10_AT | °C | Bath air temperature 10 cm below ceiling | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_bath_-10_AT | WS201809131_15 | N2_aroom_bath_-10_AT | WS201809127_12
x |  | x | x | aroom_bath_110_AT | °C | Bath air temperature @110 cm hight | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_bath_110_AT | WS201809131_03 | N2_aroom_bath_110_AT | WS201809127_15
|  | x | x | aroom_bath_110_GT | °C | Bath globe temperature @110 cm hight | not operative temperature! | Ahlborn | FPA805GTS; PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_bath_110_GT | WS201809141_09 | N2_aroom_bath_110_GT | WS201809141_17
|  | x | x | aroom_bath_110_rH | % | Bath relative air humidity @110 cm hight | - | rotronic | HC2A-S3 | IBP | 2018-08-05 | 0.8 % (rH) | 3 %rH | EL3164/EL3166 | Beckhoff | O5_aroom_bath_110_rH | 20268041 | N2_aroom_bath_110_rH | 20268016
|  | x | x | aroom_bath_170_AT | °C | Bath air temperature @170 cm hight | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_bath_170_AT | WS201809131_08 | N2_aroom_bath_170_AT | WS201809127_13
x |  | x | x | aroom_bath_heat_elP | W | Bath senible Heat Input | - | Beckhoff / MBS | EL3403‐0010 WSK 40 10/1A | IBP | 2018-12-05 | 1.5 % + 0,5% | integrated | - | - | O5_aroom_bath_heat_elP | - | N2_aroom_bath_heat_elP | -
x |  | x | x | aroom_bath_IHS_elP | W | Bath senible Internal Heat Gains set value | - | Beckhoff / MBS | EL3403‐0010 WSK 40 10/1A | IBP | 2018-12-05 | 1.5 % + 0,5% | integrated | - | - | O5_aroom_bath_IHS_elP | - | N2_aroom_bath_IHS_elP | -
|  | x | x | aroom_bed_010_AT | °C | Bed room air temperature @10 cm hight | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_bed_010_AT | WS201809131_11 | N2_aroom_bed_010_AT | WS201809127_03
|  | x | x | aroom_bed_-10_AT | °C | Bed room air temperature 10 cm below ceiling | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_bed_-10_AT | WS201809131_16 | N2_aroom_bed_-10_AT | WS201809127_01
x |  | x | x | aroom_bed_110_AT | °C | Bed room air temperature @110 cm hight | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_bed_110_AT | WS201809131_06 | N2_aroom_bed_110_AT | WS201809127_18
|  | x | x | aroom_bed_110_GT | °C | Bed room globe temperature @110 cm hight | not operative temperature! | Ahlborn | FPA805GTS; PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_bed_110_GT | WS201809141_02 | N2_aroom_bed_110_GT | WS201809141_13
|  | x | x | aroom_bed_110_rH | % | Bed room relative air humidity @110 cm hight | - | rotronic | HC2A-S3 | IBP | 2018-08-04 | 0.8 % (rH) | 3 %rH | EL3164/EL3165 | Beckhoff | O5_aroom_bed_110_rH | 20268161 | N2_aroom_bed_110_rH | 20268045
|  | x | x | aroom_bed_170_AT | °C | Bed room air temperature @170 cm hight | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_bed_170_AT | WS201809131_02 | N2_aroom_bed_170_AT | WS201809127_14
x |  | x | x | aroom_bed_heat_elP | W | Bed room senible Heat Input | - | Beckhoff / MBS | EL3403‐0010 WSK 40 10/1A | IBP | 2018-12-05 | 1.5 % + 0,5% | integrated | - | - | O5_aroom_bed_heat_elP | - | N2_aroom_bed_heat_elP | -
x |  | x | x | aroom_bed_IHS_elP | W | Bed room senible Internal Heat Gains set value | calculated from schedule and heat_elP | - | - | [6 empty] | O5_aroom_bed_IHS_elP |  | N2_aroom_bed_IHS_elP
|  | x | x | aroom_child1_010_AT | °C | Child 1 air temperature @10 cm hight | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_child1_010_AT | WS201809130_19 | N2_aroom_child1_010_AT | WS201809137_02
|  | x | x | aroom_child1_-10_AT | °C | Child 1 air temperature 10 cm below ceiling | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_child1_-10_AT | WS201809137_05 | N2_aroom_child1_-10_AT | WS201809137_12
x |  | x | x | aroom_child1_110_AT | °C | Child 1 air temperature @110 cm hight | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_child1_110_AT | WS201809130_17 | N2_aroom_child1_110_AT | WS201809137_01
|  | x | x | aroom_child1_110_GT | °C | Child 1 globe temperature @110 cm hight | not operative temperature! | Ahlborn | FPA805GTS; PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_child1_110_GT | WS201809141_05 | N2_aroom_child1_110_GT | WS201809141_19
|  | x | x | aroom_child1_110_rH | % | Child 1 relative air humidity @110 cm hight | - | rotronic | HC2A-S3 | IBP | 2018-08-06 | 0.8 % (rH) | 3 %rH | EL3164/EL3169 | Beckhoff | O5_aroom_child1_110_rH | 20268063 | N2_aroom_child1_110_rH | 20268167
|  | x | x | aroom_child1_170_AT | °C | Child 1 air temperature @170 cm hight | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_child1_170_AT | WS201809130_15 | N2_aroom_child1_170_AT | WS201809137_13
x |  | x | x | aroom_child1_heat_elP | W | Child 1 senible Heat Input | - | Beckhoff / MBS | EL3403‐0010 WSK 40 10/1A | IBP | 2018-12-05 | 1.5 % + 0,5% | integrated | - | - | O5_aroom_child1_heat_elP | - | N2_aroom_child1_heat_elP | -
x |  | x | x | aroom_child1_IHS_elP | W | Child1's senible Internal Heat Gains set value | calculated from schedule and heat_elP | - | - | [6 empty] | O5_aroom_child1_IHS_elP |  | N2_aroom_child1_IHS_elP
x |  | x | x | aroom_child1_win_pos | - | Child_1 window position | 0 is closed, 1 is open; In hourly-averaged file, this is a fraction. | Thermokon | SRW01 | - | [5 empty] | O5_aroom_child1_win_pos |  | N2_aroom_child1_win_pos
|  | x | x | aroom_child2_010_AT | °C | Child 2 air temperature @10 cm hight | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_child2_010_AT | WS201809130_19 | N2_aroom_child2_010_AT | WS201809137_04
|  | x | x | aroom_child2_-10_AT | °C | Child 2 air temperature 10 cm below ceiling | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_child2_-10_AT | WS201809131_20 | N2_aroom_child2_-10_AT | WS201809137_11
x |  | x | x | aroom_child2_110_AT | °C | Child 2 air temperature @110 cm hight | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_child2_110_AT | WS201809130_20 | N2_aroom_child2_110_AT | WS201809137_07
|  | x | x | aroom_child2_110_GT | °C | Child 2 globe temperature @110 cm hight | not operative temperature! | Ahlborn | FPA805GTS; PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_child2_110_GT | WS201809141_06 | N2_aroom_child2_110_GT | WS201809141_14
|  | x | x | aroom_child2_110_rH | % | Child 2 relative air humidity @110 cm hight | - | rotronic | HC2A-S3 | IBP | 2018-08-07 | 0.8 % (rH) | 3 %rH | EL3164/EL3170 | Beckhoff | O5_aroom_child2_110_rH | 20268061 | N2_aroom_child2_110_rH | 20267797
|  | x | x | aroom_child2_170_AT | °C | Child 2 air temperature @170 cm hight | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_child2_170_AT | WS201809130_21 | N2_aroom_child2_170_AT | WS201809137_10
x |  | x | x | aroom_child2_heat_elP | W | Child 2 senible Heat Input | - | Beckhoff / MBS | EL3403‐0010 WSK 40 10/1A | IBP | 2018-12-05 | 1.5 % + 0,5% | integrated | - | - | O5_aroom_child2_heat_elP | - | N2_aroom_child2_heat_elP | -
x |  | x | x | aroom_child2_IHS_elP | W | Child2's senible Internal Heat Gains set value | calculated from schedule and heat_elP | - | - | [6 empty] | O5_aroom_child2_IHS_elP |  | N2_aroom_child2_IHS_elP
|  | x | x | aroom_corridor_010_AT | °C | Corridor air temperature @10 cm hight | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_corridor_010_AT | WS201809130_08 | N2_aroom_corridor_010_AT | WS201809127_02
|  | x | x | aroom_corridor_-10_AT | °C | Corridor air temperature 10 cm below ceiling | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_corridor_-10_AT | WS201809131_17 | N2_aroom_corridor_-10_AT | WS201809127_07
x |  | x | x | aroom_corridor_110_AT | °C | Corridor air temperature @110 cm hight | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_corridor_110_AT | WS201809130_07 | N2_aroom_corridor_110_AT | GH201804113_18
|  | x | x | aroom_corridor_110_GT | °C | Corridor globe temperature @110 cm hight | not operative temperature! | Ahlborn | FPA805GTS; PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_corridor_110_GT | WS201809141_01 | N2_aroom_corridor_110_GT | WS201809141_16
|  | x | x | aroom_corridor_110_rH | % | Corridor relative air humidity @110 cm hight | - | rotronic | HC2A-S3 | IBP | 2018-08-06 | 0.8 % (rH) | 3 %rH | EL3164/EL3167 | Beckhoff | O5_aroom_corridor_110_rH | 2026071 | N2_aroom_corridor_110_rH | 20267929
|  | x | x | aroom_corridor_170_AT | °C | Corridor air temperature @170 cm hight | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_corridor_170_AT | WS201809130_06 | N2_aroom_corridor_170_AT | WS201809127_06
x |  | x | x | aroom_corridor_heat_elP | W | Corridor senible Heat Input | - | Beckhoff / MBS | EL3403‐0010 WSK 40 10/1A | IBP | 2018-12-05 | 1.5 % + 0,5% | integrated | - | - | O5_aroom_corridor_heat_elP | - | N2_aroom_corridor_heat_elP | -
x |  | x | x | aroom_corridor_IHS_elP | W | Corridor senible Internal Heat Gains set value | calculated from schedule and heat_elP | - | - | [6 empty] | O5_aroom_corridor_IHS_elP |  | N2_aroom_corridor_IHS_elP
|  | x | x | aroom_dining_010_AT | °C | Dining room air temperature @10 cm hight | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_dining_010_AT | WS201809130_23 | N2_aroom_dining_010_AT | WS201809131_10
|  | x | x | aroom_dining_-10_AT | °C | Dining room air temperature 10 cm below ceiling | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_dining_-10_AT | WS201809131_14 | N2_aroom_dining_-10_AT | WS201809127_10
x |  | x | x | aroom_dining_110_AT | °C | Dining room air temperature @110 cm hight | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_dining_110_AT | WS201809130_22 | N2_aroom_dining_110_AT | WS201809131_04
|  | x | x | aroom_dining_110_GT | °C | Dining room globe temperature @110 cm hight | not operative temperature! | Ahlborn | FPA805GTS; PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_dining_110_GT | WS201809141_07 | N2_aroom_dining_110_GT | WS201809141_12
|  | x | x | aroom_dining_110_rH | % | Dining room relative air humidity @110 cm hight | - | rotronic | HC2A-S3 | IBP | 2018-08-02 | 0.8 % (rH) | 3 %rH | EL3164/EL3163 | Beckhoff | O5_aroom_dining_110_rH | 20268015 | N2_aroom_dining_110_rH | 20268052
|  | x | x | aroom_dining_170_AT | °C | Dining room air temperature @170 cm hight | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_dining_170_AT | WS201809130_03 | N2_aroom_dining_170_AT | WS201809131_12
x |  | x | x | aroom_dining_heat_elP | W | Dining room senible Heat Input | - | Beckhoff / MBS | EL3403‐0010 WSK 40 10/1A | IBP | 2018-12-05 | 1.5 % + 0,5% | integrated | - | - | O5_aroom_dining_heat_elP | - | N2_aroom_dining_heat_elP | -
|  | x | x | aroom_doorway_010_AT | °C | Doorway air temperature @10 cm hight | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_doorway_010_AT | WS201809130_09 | N2_aroom_doorway_010_AT | WS201809127_08
|  | x | x | aroom_doorway_-10_AT | °C | Doorway air temperature 10 cm below ceiling | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_doorway_-10_AT | WS201809131_23 | N2_aroom_doorway_-10_AT | WS201809127_24
x |  | x | x | aroom_doorway_110_AT | °C | Doorway air temperature @110 cm hight | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_doorway_110_AT | WS201809130_10 | N2_aroom_doorway_110_AT | WS201809127_20
|  | x | x | aroom_doorway_110_GT | °C | Doorway globe temperature @110 cm hight | not operative temperature! | Ahlborn | FPA805GTS; PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_doorway_110_GT | WS201809141_10 | N2_aroom_doorway_110_GT | WS201809141_20
|  | x | x | aroom_doorway_110_rH | % | Doorway relative air humidity @110 cm hight | - | rotronic | HC2A-S3 | IBP | 2018-08-07 | 0.8 % (rH) | 3 %rH | EL3164/EL3168 | Beckhoff | O5_aroom_doorway_110_rH | 20268051 | N2_aroom_doorway_110_rH | 20268160
|  | x | x | aroom_doorway_170_AT | °C | Doorway air temperature @170 cm hight | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_doorway_170_AT | WS201809130_04 | N2_aroom_doorway_170_AT | WS201809127_17
x |  | x | x | aroom_doorway_heat_elP | W | Doorway senible Heat Input | - | Beckhoff / MBS | EL3403‐0010 WSK 40 10/1A | IBP | 2018-12-05 | 1.5 % + 0,5% | integrated | - | - | O5_aroom_doorway_heat_elP | - | N2_aroom_doorway_heat_elP | -
|  | x | x | aroom_kitchen_010_AT | °C | Kitchen air temperature @10 cm hight | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_kitchen_010_AT | WS201809130_11 | N2_aroom_kitchen_010_AT | WS201809131_22
|  | x | x | aroom_kitchen_-10_AT | °C | Kitchen air temperature 10 cm below ceiling | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_kitchen_-10_AT | WS201809131_13 | N2_aroom_kitchen_-10_AT | WS201809127_11
x |  | x | x | aroom_kitchen_110_AT | °C | Kitchen air temperature @110 cm hight | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_kitchen_110_AT | WS201809130_13 | N2_aroom_kitchen_110_AT | WS201809131_24
|  | x | x | aroom_kitchen_110_GT | °C | Kitchen globe temperature @110 cm hight | not operative temperature! | Ahlborn | FPA805GTS; PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_kitchen_110_GT | WS201809141_04 | N2_aroom_kitchen_110_GT | WS201809141_15
|  | x | x | aroom_kitchen_110_rH | % | Kitchen relative air humidity @110 cm hight | - | rotronic | HC2A-S3 | IBP | 2018-08-03 | 0.8 % (rH) | 3 %rH | EL3164/EL3164 | Beckhoff | O5_aroom_kitchen_110_rH | 20268050 | N2_aroom_kitchen_110_rH | 20268068
|  | x | x | aroom_kitchen_170_AT | °C | Kitchen air temperature @170 cm hight | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_kitchen_170_AT | WS201809130_12 | N2_aroom_kitchen_170_AT | WS201809131_18
x |  | x | x | aroom_kitchen_heat_elP | W | Kitchen senible Heat Input | - | Beckhoff / MBS | EL3403‐0010 WSK 40 10/1A | IBP | 2018-12-05 | 1.5 % + 0,5% | integrated | - | - | O5_aroom_kitchen_heat_elP | - | N2_aroom_kitchen_heat_elP | -
x |  | x | x | aroom_kitchen_IHS_elP | W | Kitchen senible Internal Heat Gains set value | - | Beckhoff / MBS | EL3403‐0010 WSK 40 10/1A | IBP | 2018-12-05 | 1.5 % + 0,5% | integrated | - | - | O5_aroom_kitchen_IHS_elP | - | N2_aroom_kitchen_IHS_elP | -
x |  | x | x | aroom_kitchen_door_pos | - | door position between kitchen and living | 0 is closed, 1 is open; In hourly-averaged file, this is a fraction. | Thermokon | SRW01 | - | [5 empty] | O5_aroom_kitchen_door_pos |  | N2_aroom_kitchen_door_pos
|  | x | x | aroom_living_010_AT | °C | Living room air temperature @10 cm hight | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_living_010_AT | WS201809130_14 | N2_aroom_living_010_AT | WS201809131_01
|  | x | x | aroom_living_-10_AT | °C | Living room air temperature 10 cm below ceiling | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_living_-10_AT | WS201809131_21 | N2_aroom_living_-10_AT | WS201809127_19
x |  | x | x | aroom_living_110_AT | °C | Living room air temperature @110 cm hight | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_living_110_AT | WS201809130_24 | N2_aroom_living_110_AT | WS201809131_09
|  | x | x | aroom_living_110_GT | °C | Living room globe temperature @110 cm hight | not operative temperature! | Ahlborn | FPA805GTS; PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_living_110_GT | WS201809141_03 | N2_aroom_living_110_GT | WS201809141_11
|  | x | x | aroom_living_110_rH | % | Living room relative air humidity @110 cm hight | - | rotronic | HC2A-S3 | IBP | 2018-08-01 | 0.8 % (rH) | 3 %rH | EL3164/EL3162 | Beckhoff | O5_aroom_living_110_rH | 20268169 | N2_aroom_living_110_rH | 20267909
|  | x | x | aroom_living_170_AT | °C | Living room air temperature @170 cm hight | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_living_170_AT | WS201809131_07 | N2_aroom_living_170_AT | WS201809131_05
x |  | x | x | aroom_living_heat_elP | W | Living room senible Heat Input | - | Beckhoff / MBS | EL3403‐0010 WSK 40 10/1A | IBP | 2018-12-05 | 1.5 % + 0,5% | integrated | - | - | O5_aroom_living_heat_elP | - | N2_aroom_living_heat_elP | -
x |  | x | x | aroom_living_IHS_elP | W | Living room senible Internal Heat Gains set value | calculated from schedule and heat_elP | - | - | [6 empty] | O5_aroom_living_IHS_elP |  | N2_aroom_living_IHS_elP
|  | x | x | aroom_socket_always_elP | W | Power consumption of socket permanent on | - | Beckhoff / MBS | EL3403‐0010 WSK 40 10/1A | IBP | 2018-12-05 | 1.5 % + 0,5% | integrated | - | - | O5_aroom_socket_always_elP | - | N2_aroom_socket_always_elP | -
|  | x | x | aroom_socket_floor_elP | W | Power consumption of socket floorwise on | - | Beckhoff / MBS | EL3403‐0010 WSK 40 10/1A | IBP | 2018-12-05 | 1.5 % + 0,5% | integrated | - | - | O5_aroom_socket_floor_elP | - | N2_aroom_socket_floor_elP | -
|  | x | x | aroom_socket_room_elP | W | Power consumption of socket roomwise on | - | Beckhoff / MBS | EL3403‐0010 WSK 40 10/1A | IBP | 2018-12-05 | 1.5 % + 0,5% | integrated | - | - | O5_aroom_socket_room_elP | - | N2_aroom_socket_room_elP | -
|  | x | x | aroom_stairs_010_AT | °C | Stairs air temperature @10 cm hight | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_stairs_010_AT | WS201809130_16 | N2_aroom_stairs_010_AT | WS201809127_23
|  | x | x | aroom_stairs_-10_AT | °C | Stairs air temperature 10 cm below ceiling | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_stairs_-10_AT | WS201809137_06 | N2_aroom_stairs_-10_AT | WS201809127_04
x |  | x | x | aroom_stairs_110_AT | °C | Stairs air temperature @110 cm hight | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_stairs_110_AT | WS201809130_01 | N2_aroom_stairs_110_AT | WS201809127_22
|  | x | x | aroom_stairs_170_AT | °C | Stairs air temperature @170 cm hight | Radiation shielded | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_aroom_stairs_170_AT | WS201809130_02 | N2_aroom_stairs_170_AT | WS201809127_21
x |  | x | x | aroom_stairs_heat_elP | W | Stairs senible Heat Input | - | Beckhoff / MBS | EL3403‐0010 WSK 40 10/1A | IBP | 2018-12-05 | 1.5 % + 0,5% | integrated | - | - | O5_aroom_stairs_heat_elP | - | N2_aroom_stairs_heat_elP | -
x |  | x | x | cellar_285_AT | °C | Air temperature 10 cm under the cellar's ceiling |  | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_cellar_285_AT | WS201802101_01 | N2_cellar_285_AT | WS201809138_04
x |  | x | x | cellar_285_AT2 | °C | Air temperature 10 cm under the cellar's ceiling |  | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K (±1.5K under full sol. irradiation) | 0.1 K | EL3202-0010 | Beckhoff | O5_cellar_285_AT2 | WS201809137_22 | N2_cellar_285_AT2 | WS201809137_23
x |  | x | x | downwards_SW_rad | W/m² | downwards solar radiation in front of living room for albedo calculation |  | Kipp & Zonen | CMP11 | Kipp & Zonen | 2018 | < 2 % | ± 7,5 W/m² | EL3602-0010 | Beckhoff | O5_downwards_SW_rad | 101656 | N2_downwards_SW_rad | 80534
|  |  |  | hp_status_command |  | on/off command applied to heat pump
|  |  |  | hp_supply_temp_command |  | supply water temperature command applied to heat pump
|  | x |  | hp_el_cons | °C | Heat Pump electricity consumption | - | Thermosensor 1/10 DIN B | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K | 0.1 K | EL3202-0010 | Beckhoff | O5_hp_el_cons | TT201808123_20 |  | -
x |  | p |  | HeatPump_Tsup | °C | Actual Heat Pump supply temperature | - | Thermosensor 1/10 DIN B | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K | 0.1 K | EL3202-0010 | Beckhoff | O5_HeatPump_Tsup | TT201808123_12 |  | -
|  | x | x | kitchen_door_pos | - | position of kitchen door | Wireless - EnOcean | Thermokon | SRW01 | - | - | - | - | KL6581/6583 | Beckhoff | O5_kitchen_door_pos |  | N2_kitchen_door_pos
|  | x |  | Vent_absH_IMS | g/kg | absolute humidity of internal moisture source | calculated | - | - | - | - | - | - | - | - | O5_Vent_absH_IMS | -
|  | x |  | Vent_absH_IMS_soll | g/kg | set value of absolute humidity | Set Value | - | - | - | - | - | - | - | - | O5_Vent_absH_IMS_soll | -
|  | x |  | Vent_absH_tot | g/kg | absolute humidity of supply air | calculated | - | - | - | - | - | - | - | - | O5_Vent_absH_tot | -
|  | x |  | Vent_absH_treq | °C | set temperature of internal moisture source | Set Value | - | - | - | - | - | - | - | - | O5_Vent_absH_treq | -
|  | x |  | Vent_afterIHS_AT | °C | Ventilation air temperature after heating | - | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K | 0.1 K | EL3202-0010 | Beckhoff | O5_Vent_afterIHS_AT | WS201809138_08
|  | x |  | Vent_afterIMS_rH | % | Relative Humidity ofter the Moisture source |  | Rotronic | HC2A-S3 | IBP | 2018-08-07 | 0.8 % (rH) | 3 %rH | EL3164/EL3167 | Beckhoff | O5_Vent_afterIMS_rH | 20268014
|  | x | x | Vent_bath_EHA_AT | °C | Exhaust air temperature bath | - | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K | 0.1 K | EL3202-0010 | Beckhoff | O5_Vent_bath_EHA_AT | WS201809138_23 | N2_Vent_bath_EHA_AT | WS201809137_18
x |  | x | x | Vent_bath_EHA_VFR | m³/h | Exhaust air volume flow rate bath | - | Schmidt Technology GmbH | SS 20.501 | Electro-Mation Ing. Büro GmbH | 2018-07-03 | ±3.5 m³/h | 0.03% | EL3164 | Beckhoff | O5_Vent_bath_EHA_VFR | 205002642 | N2_Vent_bath_EHA_VFR | 205002638
x |  | x | x | Vent_child1_child2_EHA_fan_elP | W | Child 1/child 2 power consumption of the exhaust air fan |  | Beckhoff | EL3403‐0010 | IBP | 2018-12-05 | 1.5 % | integrated | - | - | O5_Vent_child1_child2_EHA_fan_elP | - | N2_Vent_child1_child2_EHA_fan_elP | -
x |  | x | x | Vent_child1_child2_SUA_fan_elP | W | Child 1/child 2 power consumption of the supply air fan |  | Beckhoff | EL3403‐0010 | IBP | 2018-12-05 | 1.5 % | integrated | - | - | O5_Vent_child1_child2_SUA_fan_elP | - | N2_Vent_child1_child2_SUA_fan_elP | -
x |  | x | x | Vent_child1_child2_SUA_IHS_elP | W | Child 1/child 2 power consumption of the supply air heater |  | Beckhoff / MBS | EL3403‐0010 WSK 40 10/1A | IBP | 2018-12-05 | 1.5 % + 0,5% | integrated | - | - | O5_Vent_child1_child2_SUA_IHS_elP | - | N2_Vent_child1_child2_SUA_IHS_elP | -
|  | x | x | Vent_child1_EHA_AT | °C | Exhaust air temperature child 1 | - | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K | 0.1 K | EL3202-0010 | Beckhoff | O5_Vent_child1_EHA_AT | WS201809138_16 | N2_Vent_child1_EHA_AT | WS201809138_07
x |  | x | x | Vent_child1_EHA_VFR | m³/h | Exhaust air volume flow rate child 1 | - | Schmidt Technology GmbH | SS 20.501 | Electro-Mation Ing. Büro GmbH | 2018-07-05 | ±3.5 m³/h | 0.03% | EL3164 | Beckhoff | O5_Vent_child1_EHA_VFR | 205000785 | N2_Vent_child1_EHA_VFR | 205000839
x |  | x | x | Vent_child1_SUA_AT | °C | Supply air temperature child 1 | - | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K | 0.1 K | EL3202-0010 | Beckhoff | O5_Vent_child1_SUA_AT | WS201809138_05 | N2_Vent_child1_SUA_AT | WS201809138_09
x |  | x | x | Vent_child1_SUA_VFR | m³/h | Supply air volume flow rate child 1 | - | Schmidt Technology GmbH | SS 20.501 | Electro-Mation Ing. Büro GmbH | 2018-07-04 | ±3.5 m³/h | 0.03% | EL3164 | Beckhoff | O5_Vent_child1_SUA_VFR | 205002640 | N2_Vent_child1_SUA_VFR | 205002641
|  | x | x | Vent_child2_EHA_AT | °C | Exhaust air temperature child 2 | - | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K | 0.1 K | EL3202-0010 | Beckhoff | O5_Vent_child2_EHA_AT | WS201809138_13 | N2_Vent_child2_EHA_AT | WS201809138_17
x |  | x | x | Vent_child2_EHA_VFR | m³/h | Exhaust air volume flow rate child 2 | - | Schmidt Technology GmbH | SS 20.501 | Electro-Mation Ing. Büro GmbH | 2018-07-07 | ±3.5 m³/h | 0.03% | EL3164 | Beckhoff | O5_Vent_child2_EHA_VFR | 205002619 | N2_Vent_child2_EHA_VFR | 205002581
x |  | x | x | Vent_child2_SUA_AT | °C | Supply air temperature child 2 | - | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K | 0.1 K | EL3202-0010 | Beckhoff | O5_Vent_child2_SUA_AT | WS201809138_14 | N2_Vent_child2_SUA_AT | WS201809138_03
x |  | x | x | Vent_child2_SUA_VFR | m³/h | Supply air volume flow rate child 2 | - | Schmidt Technology GmbH | SS 20.501 | Electro-Mation Ing. Büro GmbH | 2018-07-06 | ±3.5 m³/h | 0.03% | EL3164 | Beckhoff | O5_Vent_child2_SUA_VFR | 205002621 | N2_Vent_child2_SUA_VFR | 205002620
|  | x |  | Vent_CO2amb | ppm | CO2 concentration of ambient air | In outside air duct | Vaisala | GMP222 | Vaisala |  | ± (1.5 % of range + 2 % of reading) |  |  | Beckhoff | O5_Vent_CO2amb | F2840051
x |  | x | x | Vent_dining_bath_EHA_fan_elP | W | Dining/baths' power consumption of the exhaust air fan |  | Beckhoff | EL3403‐0010 | IBP | 2018-12-05 | 1.5 % | integrated | - | - | O5_Vent_dining_bath_EHA_fan_elP | - | N2_Vent_dining_bath_EHA_fan_elP | -
|  | x | x | Vent_dining_EHA_AT | °C | Exhaust air temperature dining room | - | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K | 0.1 K | EL3202-0010 | Beckhoff | O5_Vent_dining_EHA_AT | WS201809138_19 | N2_Vent_dining_EHA_AT | WS201809137_19
x |  | x | x | Vent_dining_EHA_VFR | m³/h | Exhaust air volume flow rate dining room | - | Schmidt Technology GmbH | SS 20.501 | Electro-Mation Ing. Büro GmbH | 2018-07-02 | ±3.5 m³/h | 0.03% | EL3164 | Beckhoff | O5_Vent_dining_EHA_VFR | 205002624 | N2_Vent_dining_EHA_VFR | 205002625
|  | x | x | Vent_FL_child1_calc_thP | W | thermal power added/removed by mechanical ventilation in child 1 | calculated | - | - | - | - | - | - | - | - | O5_Vent_FL_child1_calc_thP | - | N2_Vent_FL_child1_calc_thP | -
|  | x | x | Vent_FL_child2_calc_thP | W | thermal power added/removed by mechanical ventilation in child 2 | calculated | - | - | - | - | - | - | - | - | O5_Vent_FL_child2_calc_thP | - | N2_Vent_FL_child2_calc_thP | -
|  | x | x | Vent_GF_calc_thP | W | thermal power added/removed by mechanical ventilation in ground floor | calculated | - | - | - | - | - | - | - | - | O5_Vent_GF_calc_thP | - | N2_Vent_GF_calc_thP | -
x |  | x | x | Vent_living_SUA_AT | °C | Supply air temperature living room | - | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K | 0.1 K | EL3202-0010 | Beckhoff | O5_Vent_living_SUA_AT | WS201809138_24 | N2_Vent_living_SUA_AT | WS201809137_14
x |  | x | x | Vent_living_SUA_fan_elP | W | Living rooms' power consumption of the supply air fan |  | Beckhoff | EL3403‐0010 | IBP | 2018-12-05 | 1.5 % | integrated | - | - | O5_Vent_living_SUA_fan_elP | - | N2_Vent_living_SUA_fan_elP | -
x |  | x | x | Vent_living_SUA_IHS_elP | W | Living rooms' power consumption of the supply air heater |  | Beckhoff / MBS | EL3403‐0010 WSK 40 10/1A | IBP | 2018-12-05 | 1.5 % + 0,5% | integrated | - | - | O5_Vent_living_SUA_IHS_elP | - | N2_Vent_living_SUA_IHS_elP | -
x |  | x | x | Vent_living_SUA_VFR | m³/h | Supply air volume flow rate living room | - | Schmidt Technology GmbH | SS 20.501 | Electro-Mation Ing. Büro GmbH | 2018-07-01 | ±3.5 m³/h | 0.03% | EL3164 | Beckhoff | O5_Vent_living_SUA_VFR | 205002644 | N2_Vent_living_SUA_VFR | 205002643
|  | x |  | Vent_out_absH_amb | g/kg | absolute humidity of ambient air | calculated | - | - | - | - | - | - | - | - | O5_Vent_out_absH_amb | -
|  | x |  | Vent_Tamb | °C | temperature of ambient air | - | Heraeus Holding GmbH | PT100; 4-wire | IBP calibration lab | See certificate no. | ±0.015K | 0.1 K | EL3202-0010 | Beckhoff | O5_Vent_Tamb | WS201809138_22
|  | x | x | Window | - | position of operable external window | Wireless - EnOcean | Thermokon | SRW01 | - | - | - | - | KL6581/6583 | Beckhoff | O5_Window |  | N2_Window